Operation and Determination of the          Complex Permittivity

Permittivity Meter

Name		Teichtmeister Marcus

Date		06.12.2025

Table of Contents

1	Operational Amplifier	iii

2	Operation	1

2.1	Sine Wave Filter	1

2.2	Pre-Amplifier	4

3	6

# Operational Amplifier

There are three operational amplifiers to choose from: OPA836 and OPA2690 from Texas Instruments and ADA4807 from Analog Devices.

To be able to choose which operational amplifier suits best, a simulation in closed loop has been made. The amplification |Gjω| in passband is about 20dB.

Figure ‎1-1	Non-Inverting Amplifier using ADA4807

Figure ‎1-2	Bodeplot of Different Operational Amplifiers

Figure ‎1-2 shows, that OPA2690 has a high amplification at 20MHz, therefore it is going to be used in the amplifier stage. Besides that, OPA836 has a weaker amplification than ADA4807, OPA836 will be used for pre-amplifier and as a voltage follower because a higher voltage is not needed for the notch filter and due to costs.

Table ‎1-1	Amplification and Phase Shift at 19,999MHz

|  | argGjω in ° | Gjω in dB |
| --- | --- | --- |
| ADA4807 | -85,686 | 15,133 |
| OPA836 | -83,330 | 7,7489 |
| OPA2690 | -38,023 | 19,522 |

# Operation

## Sine Wave Filter

The sine wave filter 7th order has been replaced by 9th. Reason for that is decrease of damping higher harmonics. In addition to that, the 7th sine wave filter from Rippel is using 39Ω termination. This gives a mismatch both when using the network analyser and in the simulation. Improvements in terms of matching impedance and filtering higher harmonics have been made.

It is important to mention that the input voltage is not allowed to be greater than 2V, because it will lead to distortion and the amplifier stage will not be able to perform correctly.

Sine Wave Filter is used to filter the harmonics of the input rectangle wave, resulting in having a sinusoidal wave for the notch filter.

Figure ‎2-1	Sine Wave Filter 9th Order

A voltage follower is used to not load the DAC-output of the microcontroller. Resistances which are in series to the inductances represents the serial resistance of a real inductor.

Figure ‎2-2	Smith Chart of Sine Wave Filter

Figure ‎2-2 shows a great mismatch of the 7th order lowpass filter. The difference between the ideal and real 9th order lowpass filter is, that the ideal one is from RF-Design and the real one has higher series resistance.

Figure ‎2-3	Bodeplot of Sine Wave Filter

Figure ‎2-3 shows that the 9th order decreases faster than the 7th order. In addition to that it has a small ripple. Reason for that is the use of a Chebyshev filter, which is used due to its high falling edge.

Figure ‎2-4	Pol-Zero-Map of 9th Order Sine Wave Filter

As seen in Figure ‎2-4, all poles are lying on the Chebyshev-Ellipse, which is a sign of a Chebyshev filter.

## Pre-Amplifier

Besides the use of a different operational amplifier, the topology of the pre-amplifier has not changed.

Figure ‎2-5	Pre-Amplifier

Figure ‎2-6	Bodeplot of Pre-Amplifier

Table ‎2-1	Amplification and Phase Shift at 20,045MHz

| ¦Gjω¦ in dB | arg[Gjω] in ° |
| --- | --- |
| 4,2966 | -41,95 |

## Notch Filter

This is the main part of the permittivity meter. It has its resonance frequency fres at 20MHz, at this frequency the output should be ideally 0V. To move the minimum along the frequency, the varactor-diode D1 needs to be tuned. Varactor-diode D2 is used to increase the quality factor.

During calibration, the sensor is held at a reference medium, air. Both varactor-diodes will be tuned to have its minimum at 20MHz and a high-quality factor. When the sensor is put in a different medium, e.g. snow, then the resonance frequency will be smaller due to the additional capacitance of snow ΔC and because of conductivity losses, which is represented in the schematic as a fictional resistance Rp, the quality factor Q will also change.

Both parts a necessary to determine the complex permittivity of snow. The calculation of the complex permittivity can be seen in ‎3.

Figure	‎2-7Notch Filter

Figure ‎2-8	Bodeplot of Notch-Filter when UD1 varies

Figure ‎2-9	Bodeplot of Notch-Filter when UD2 varies

Figure ‎2-10	Bodeplot of Notch-Filter when CS varies

Figure ‎2-11	Bodeplot of Notch-Filter when Rp varies

## Amplifier Stage

The only difference in the amplifier stage is the use of a difference transistor. BFP420 BJT NPN transistor from Infineon will be used. The main advantage is its small capacitance (>1pF).

Figure ‎2-12	Amplifier Stage

The real transistor from Infineon has two emitter pins, which is normal for RF-transistors. Reason for that is the reduction of parasitic inductance.

Figure ‎2-13	Bodeplot of Amplifier Stage

Figure ‎2-13 shows, that when the transistor opens there are hardly any differences at 20MHz. However, when the transistor is closed, there is a difference between the ideal and real circuit at 20MHz. Reason for that might be the small signal parameters of the transistor.

## Voltage Follower

Voltage Follower will be used, in order to measure the output voltage with the ADC from the microcontroller.

Figure ‎2-14	Voltage Follower

Figure ‎2-15	Bodeplot Voltage Follower

# Determination of the Complex Permittivity

## Formulars

The definition of the complex permittivity is:

εc=ε'-jε''=εr-jσω

With the complex permittivity the admittance of the sensor capacitance can be derived:

Y=jωε0εck=jωC0ε'-jε''=jωε'C0+ωε''C0

IY=jωε'C0=jωC0+ΔC→ε'=1+ΔCC0

Where ΔC is the difference between the capacitance in a reference medium, e.g. air, C0 and the capacitance in a measured medium, e.g. snow, C.

RY=ωε''C0=1Rp→ε''=1ωRpC0=1ωkε0Rp

With k is the sensor specific constant k=C0ε0.

Real part of the complex permittivity can also calculate with the empirical formula from Denoth.

ε'=1+kDlogVMVRef

Where VM is the varactor-diode voltage UD1 to move the resonance frequency back to 20MHz, when the sensor is in a medium. VRef is the needed voltage UD1 to have the resonance frequency at 20MHz when the sensor is in air. kD is the slope of the varactor-capacitance in respect to the voltage relative to the sensor capacitance.

kD=C2-C1C0logV2V1

## Calculation of the Relative Permittivity

kD of BB353 from Infineon, C1, C2, V1 and V2 were obtained from Dierer’s thesis, page 53:

kD=C2-C1CslogV2V1=19,37-15,7820log21=0,5963

For the simulation the sensor capacitance Cs will be increased to 21pF, therefore ΔC=1pF.

εexpected'=1+ΔCCs=1+120=1,05

For VM and VRef the simulation gives 0,73V and 0,56V.

εcalculated'=1+klogVMVRef=1+0,5963log0,730,56=1,0687

This gives a relative error of:

r=εcalculated'-εexpected'εexpected'=1,0687-1,051,05=0,0178≜1,78%